## Slide 1: 声学膜片阻抗测试系统
Theory Introduction
技术支持： 18362654625（微信同号） 王博士

## Slide 2
10kHz tube, inner diameter 20mm.

## Slide 3
10kHz tube, inner diameter 20mm.

## Slide 4
10k Hz tube, inner diameter 20mm.
Sound Source
Sample under Test
Mic 3
Mic 2
Mic 1
Mic 5(Phantom Mic)
Mic 4
-x1
-x2
-x3
0
L
0+
rs
r
-Ls
Rigid Wall
将rs反射波,与第一次入射波叠加，计算r

## Slide 5
Mic Distance | Frequency Range
80mm(x1-x3) | 214Hz - 1929Hz
15mm(x2-x3) | 1143Hz - 10290Hz
Tube Effective range | Mics on work
10Hz - 220Hz | Mic 3; Mic 4
221Hz - 1500Hz | Mic 1; Mic 3; Mic 4
1501Hz - 10kHz | Mic 2; Mic 3; Mic 4
Microphone Layout:
coordinate | value
x1 | 87mm
x2 | 22mm
x3 | 7mm
低频P0- = P3 中高频通过P3，P1计算r，再计算P0-

## Slide 6
 | Tube Effective range | Mics on work
Low-frequency | 10Hz - 220Hz | Mic 3; Mic 4
Mid-range | 221Hz - 1500Hz | Mic 1; Mic 3; Mic 4
High-frequency | 1501Hz - 10kHz | Mic 2; Mic 3; Mic 4
Low - frequency
Mid-range & high - frequency
A为左管Pia
计算p0和p3时，都已经计算了rs的反射，相除rs被约掉了
中频 高频 计算 r，用到的麦克风不一样

## Slide 7
Mid-range:
high - frequency:
 | Tube Effective range | Mics on work
Low-frequency | 10Hz - 220Hz | Mic 3; Mic 4
Mid-range | 221Hz - 1500Hz | Mic 1; Mic 3; Mic 4
High-frequency | 1501Hz - 10kHz | Mic 2; Mic 3; Mic 4
r

## Slide 8
Low - frequency
Mid-range & high - frequency

## Slide 9
20kHz tube, inner diameter 10mm.

## Slide 10
20kHz tube, inner diameter 10mm.

## Slide 11
20kHz tube, inner diameter 10mm.
Sound Source
Sample under Test
Mic 3
Mic 2
Mic 1
Mic 4
-x1
-x2
-x3
0
Ls
x3
rs
r1
-Ls
Absorption terminal
r2
r3
Input Tube
Transmission Tube

## Slide 12
Mic Distance | Frequency Range
50mm | 343Hz - 3087Hz
7.5mm | 2286Hz - 20580Hz
Tube Effective range | Mics on work
500Hz - 2500Hz | Mic 1; Mic 3; Mic 4
2501Hz - 20kHz | Mic 2; Mic 3; Mic 4
Microphone Layout:
coordinate | value
x1 | 57mm
x2 | 14.5mm
x3 | 7mm

## Slide 13
 | Tube Effective range | Mics on work
Mid-range | 500Hz - 2500Hz | Mic 1; Mic 3; Mic 4
High-frequency | 2501Hz - 20kHz | Mic 2; Mic 3; Mic 4
Mid-range & high - frequency
Input Tube

## Slide 14
 | Tube Effective range | Mics on work
Mid-range | 500Hz - 2500Hz | Mic 1; Mic 3; Mic 4
High-frequency | 2501Hz - 20kHz | Mic 2; Mic 3; Mic 4
Mid-range & high - frequency
Transmission Tube
r3与r1等价
r2与rs等价

## Slide 15
Mid-range:
high - frequency:
r3
 | Tube Effective range | Mics on work
Mid-range | 500Hz - 2500Hz | Mic 1; Mic 3; Mic 4
High-frequency | 2501Hz - 20kHz | Mic 2; Mic 3; Mic 4

## Slide 16
Transmission Tube
Input Tube

## Slide 17
Q&A